## GLOBAL SPECIALTIES

INSTRUMENTS

# 1523

2.6KW Autoscan True RMS

Power Meter

### INSTRUCTION  MANUAL

TABEL OF CONTENTS

SECTION	DESCRIPTION	PAGE NO.

1	SAFETY INSTRUCTIONS	1

2	TECHNICAL SPECIFICATIONS	2

3	GENERAL DESCRIPTION	3

4	OPERATING INSTRUCTION	4

5	REMOTE COMMUNICATION	6

6	TESTING & CALIBRATION PROCEDURE	8

7	PART LIST	11

8	CIRCUIT DIAGRAMS & SCHEMATICS	16

SECTION - 1

#### SAFETY INSTRUCTIONS

The general safety information in this part of the summary is for both operating and servicing personnel.  Specific warnings and cautions will be found throughout the manual where they apply, but may not appear in this summary.

1.	  This indicates refer to the manual wherever necessary.

2.	E /    Indicates Protective Ground Terminal.

3.	  Indicates Danger ! High Voltage.

4.	Danger Arising from Loss of Ground

Upon loss of the protective-ground connection, all accessible conductive parts (including knobs and controls that may appear to be insulating) can render an electric shock.

5.	Use the Proper Fuse

To avoid fire hazard, use only the fuse of correct type, voltage rating and current rating as specific in the parts list for your product.  Refer fuse replacement to qualified service personnel.

6.	Grounding the Product

This product is grounded through the grounding conductor of the power cord.  To avoid electrical shock, plug the power cord into a properly wired receptacle before connecting to the product input or output terminals.  A protective ground connection by way of the grounding conductor in the power cord is essential for safe operation.

SECTION - 2

TECHNICAL SPECIFICATIONS

2.1	Measuring Parameters	:	True RMS V, A, VA, W and Frequency.

2.2	Measuring Limits

a)	Voltage	:	75V to 130V (115V operation).

b)	Current	:	0.2 to 20A rms AC (40A peak).

c)	Volt Amperes	:	16 to 2600VA.

d)	Power	:	1 to 2600W.

e)	Power Factor	:	0.5 to unity (No display).

f)	Frequency	:	45 to 65Hz.

2.3	Measurement Accuracy

a)	Voltage / Current	:	0.5% of reading ±1 digit.

b)	Volt Amperes/Power:	1% of reading ±2 digits.

c)	Power Factor	:	2% of reading ±2 digits.

d)	Frequency	:	0.25% of reading ±1 digit.

Note	:	i)	Accuracy valid only between above specified measurement limits.

ii)	For VA, Power  accuracies are valid only when voltage and current are between their respective limits.

2.4	Measurement Resolution

a)	Voltage	:	0.1V.

b)	Current 	:	1mA  upto 1 Amp., 10 mA above 1Amp.

c)	Volt Amperes	:	1VA.

d)	Power	:	1 Watt.

e)	Frequency	:	0.01Hz.

2.5	Miscellaneous

a)	Display	:	16x2 Alphanumeric LCD display.

e)	Interface	:	RS232.

2.6	Power	:	75V to 130V (115V operation).

2.7	Instrument Power	:	9W approx.

Consumption

2.8	Operating Temperature	:  5°C to 50°C, 90% RH.

2.9	Dimensions (mm)	:	155 (W) x 190 (H) x 286 (D) approx.

2.10	Weight	:	3 Kg. approx.

SECTION - 3

#### GENERAL DESCRIPTION

  - INTRODUCTION

Global Specialities' 2.6KW AUTOSCAN TRUE RMS POWER METER, MODEL 1523 is a state-of-the-art instrument, ideal for providing all the information about any electrical load connected to the AC Mains.  It  monitors/scans sequentially all electrical parameters viz., Voltage, Current, Power, VA, & Mains Frequency, thus providing a clear picture of the load.

User operation is very simple.  The load which is to be monitored is connected to the ‘Output’ Socket provided on the rear panel of the 1523  instrument.  Each output is rated upto 130V & 10A rms current.  The instrument is then connected to the 115V AC mains.  All the parameters are then displayed on the LCD on front panel of the instrument.

An optional facility of remote communication has been provided.  Through this, an external PC can control the instrument and obtain information from it.  This is possible by making use of a set of commands, details of which are explained later.

  - RECEIVING INSTRUCTIONS

The instrument is packed in a corrugated box.  For unpacking, remove the box cover & packing material.  Then bring out the instrument from the box.

Check carefully, whether there is any physical damage to the instrument during transit.  Inform the carrier about the damage alongwith the serial number of the instrument (located on the bottom cover).

  - POWER REQUIREMENTS

The wiring of the instrument is done as per the marking on the rear panel, viz ‘Mains In’ 105V - 130V AC, 50/60Hz.

  - INSTALLATION

The instrument is ready for operation when shipped from the factory.  The appliance/load whose power is to be monitored should be at a safe distance from the instrument.  This prevents unnecessary heating up of the instrument.

SECTION - 4

#### OPERATING INSTRUCTION

4.1	IDENTIFICATION OF CONTROLS

1.	V   -	This key is to display the measured mains rms voltage in Volts on  LCD.

2.	W  -	This key is to display the power consumed by load, every second, on LCD.

3.	F    -	This key is to display the mains frequncy in Hz on LCD.

4.	I / RANGE -  This key is to display the rms value of measured load current in Amps on LCD.

5.	VA  -  This key is to display the volt-amp of  load on LCD and it is updated every second.

6.	AUTO SCAN / GL - If instrument is not in remote mode then this key acts as entry into auto scan mode. All parameters ie V,I,W,F,VA  are scanned sequetially  with a period of 1 second.

If any other key is pressed in auto scan mode,then unit comes out of auto scan mode and the selected parameter is displayed.

If the unit is in remote mode of operation,this key acts as Go Local key.It brings the unit into local mode of operation.

7.	DISPLAY- This is 16 x 2 Alphanumeric LCD display with green backlight.

- Mains In -  Cord for delivering Mains Power(115V ac) to the instrument.

- Output - Socket for plugging in the appliance / load to be monitored.There are 2 socets to share the load of 20A equally.

- Load Switch  - Turns ON/OFF outgoing power to the load.Their are 2 switches for 2 output sockets.

- Fuse - Protective Fuse between mains & load.

- RS232 Interface Connector -  External PC communicates to the instrument using this connector.

4.2	ELECTRICAL CONNECTIONS & MEASUREMENT

Follow the diagram & the instructions given below  :

Fig. 3  Electrical connectons

a)	Switch OFF the equipment under test.

b)	Connect the Mains plug of the equipment/appliance, whose power is to be measured into one of the Output sockets (2) of the 1523

c)	With the corresponding load  switch (3) of the 5051A in the ‘OFF’ position, Turn the mains power on.

e)	Switch ON the equipment/appliance & switch on the supply to the load with the corresponding load switch(3).

Note : It is recommended that the 1523 be first switched ON and then the equipment under test.

4.3	METER OPERATION

The ‘LCD display’ (7) on front panel shows the measured parameter selected. 	On power ‘ON’, the  Manual Scan mode is chosen. In this mode parameter selected with front panel key displayed constantly on display.

- Note : If the user connects some load to the instrument & switches it on with help of load switch(3) and if the current displayed by the instrument is 0.000A then it indicates a blown fuse.  User should replace it with a new fuse.

SECTION - 5

#### REMOTE COMMUNICATION

5.1	INTRODUCTION

By using the remote communication facility and the set of remote commands, a PC can completely control the instrument and get information from it.  During remote communication, front panel is disabled (except Go Local key) as long as the remote PC doesn’t relinquish control of the instrument or the GO Local key on front panel is pressed.  Follow the diagram below for connections.

Fig. 4  Serial communication Connections

a)	Connect  the com port of the PC to the female 9 pin D-type ‘RS232’ interface connector (5) on the rear panel of the instrument.

b)	Run any communication software like hyper terminal on the PC to communicate to the instrument.Communication protocols are described as under.

5.2	COMMANDS FOR COMMUNICATION

Note :  1.Set Baud Rate = 9600 bits per second for serial communication.

Command Set:-

1.	RV - Set ‘V’ function & Read input voltage in volts.

e.g. XXX.X ‘V’

2.	RI - Set ‘I’ function & Read load current in amperes.

e.g. X.XXX ‘A’ or XX.XX ‘A’

3.	RP - Set ‘P’ function & Read apparent power in Volt-Amperes.

e.g. XXXX ‘VA’

4.	RW - Set ‘W’ function & Read Actual power in Watts.

e.g. XXXX ‘W’

5.	RF - Set ‘F’ function & Read mains frequency in Hz.

e.g. XX.XX ‘Hz’

6.	GL - To come out of Remote operation (i.e. Local Mode). This command is acknowledged by an answerback character. i.e. ’Y’.

7.	SA -  To enter into auto scan mode of all parameters. This command is acknowledged by an answerback character. i.e. ’Y’.

8.	Z - To flush buffer.

Note : To come out of Remote mode, GL switch on front panel can also be used.

5.3	NOTES REGARDING COMMUNICATION

i)	Definition of signals on ‘RS232’ interface connector on rear panel (5) for remote communication.

Pin No.	Signal	Source	Description

2	RXD	Remote PC	Receive data from Remote PC

3	TXD	Instrument	Transmit data to Remote PC

5	Gnd	—	Signal Gnd

ii)	Communication parameters

Bits per second	-	9600

Data bits	-	8

Parity	-	None

Stop bits	-	1

Flow Control	-	None

- The commands to the instrument should be in upper case letters only.

- The instrument will not recognise if any wrong command is sent.  Use flush buffer command before proceeding.

SECTION - 6

#### TESTING & CALIBRATION PROCEDURE

6.1	TITLE :	Testing & Calibration Procedure of Power Meter 1523

6.2	REF. DOCUMENTS  :    5050Power Meter Test Procedure.

6.3	EQIPMENTS :	1)	Power Meter

2)	Voltmeter for AC Voltage & Current Meter up to 20A true RMS type.

3)	Variac sourcing current up to 20A.

4)	Frequency Counter.

5)	Load up to 20A AC.

6.4	PROCEDURE  :

6.4.1	Testing of PCBs :

1.	Check all the PCBs for dry solder, loose contact of components capacitor polarities & ratings etc.

2.	Check for proper contacts of smd resistor networks of 10Kohms, on Main Board.

3.	Connect all the Boards as per wiring diagrams.

6.4.2	Testing of Power Supply :

1.	Switch on the Mains supply ( 115V ) to the unit.

2.	Check the various  supply  voltages  on CPU Board such as +5V, -5V  at  test  points  TP3 & TP4 wrt GND & reference voltage with U1-12 should be 2.500V wrt GND.

6.4.3	Power on Display :

Switch on mains. On power on welcome message appears which is as follows :

APLAB MODEL	:	1523

TRUE RMS MAINS	:	POWER METER

VERSION		:	V 1.00

Now  the system carries out eprom-test & initializes the system variables if necessary. On Successful EPROM test is shown.

EPROM TEST	:	OK

Else the following message is displayed :

EPROM TEST	:	FAIL

V  	:    ININ	W :	IGNOR

If V key is pressed it initializes the eprom. If the system is unable to initialize the eprom, it displays following message on 2nd line of display:

EPROM ERROR  :	Else if W key is pressed it ignores EPROM fail. Now Mains voltage screen is displayed on LCD.

6.4.4	Testing Microcontroller 859S52 :

1.	Check for momentary high ( +5V ) pulse at pin 9 (RESET) of 89S52 on power on or pressing reset switch SW1 on CPU board.

2.	Check pin 30 (ALE) of U1 (89S52) to be a pulse signal of 2MHz with counter.

3.	Check frequency at TP1 wrt GND to be 2MHz with counter.

4.	Keep all other presets & trimmers in middle position. Adjust preset VR1 on CPU board for proper contrast of LCD.

6.4.5	Testing of Manual & Auto Mode :

1.	All the parameters can be viewed with the keys on front panel i.e. V, I, W, VA, F.

2.	Compare V & I values on display with standard DMM. Also calculate & compareVA & Power for ( for resistive load power should be equal to VA ).

3.	Auto scanning of all parameters can be done by pressing Auto scan / GL key. Auto scanning is done at speed of 1 Sec. To come out of auto scan mode, press any other key.

6.4.6	Mains Voltage Calibration :

Set mains input voltage to the unit is at 115.0 Volts at 50/60Hz frequency and monitor voltages at DMM & Power meter simultaneously. Adjust the display to be the same with help of preset VR1 on main board. Check voltages for the range specified ( 75V - 130V ).

6.4.7	Mains Current Calibration :

1.	Connect 500mA of load to the system & adjust 0.500A on power meter display with the help of preset VR2.

Now to calibrate the current throughout the range up to 20A, following procedure is to be followed.

2.	Connect 250mA load physically & switch off the load switch on the meter.

3.	To enter software calibration mode, reset the system & press Fkey. After successful eprom test, systems shows following screen.

CALIBRATION ?

YES =  PRESS W KEY

NO  =  PRES VA KEY

4.	If VA key is pressed then the system exists calibration mode & mains voltage in manual mode is displayed.

5.	To go into calibration mode, connect port pin 1.5 ( pin 5 of U2-89S52 ) to ground & press W key.

6.	Go to current mode by pressing I switch. ‘CAL’ message will be displayed along with normal manual current mode screen.

7.	W & VA keys will now function to increase & decrease the display counts respectively.

8.	Use standard DMM to get same current on standard meter and on power meter.

9.	Press F & AUTO SCAN/GL keys simultaneously to save the calibrated current value & Exit Software Calibration Mode. The message ‘CALIBRATION SAVED’  will be flashed to confirm saving.

10.	Follow steps from 2 to 9 to calibrate the system on loads of 750mA, 5A & 15A & then verify all readings.

11.	Software Calibration mode will exit automatically without saving changes after 25 sec of idle operation, it will show following message on display :

CALIBRATION

##### TIMEOUT

SECTION - 7

#### PART LIST

PCB Components	5051A-MAIN-0306 IRD 551

=======================================================================

Ref Designator		Value

========================================================================

RESISTORS

R1			470E SMD 0805 CHIP

R2			470E SMD 0805 CHIP

R3			16E CHIP 1/8W  1206

R4			1K  1% 0805 0.125W SMD

R5			10E SMD 0805 CHIP

R6			10K SMD 0805 CHIP

R7			12K SMD 0805 CHIP

R8			470E SMD 0805 CHIP

R9			470E SMD 0805 CHIP

R10			1K  1% 0805 0.125W SMD

R11			1K  1% 0805 0.125W SMD

R12			10K SMD 0805 CHIP

R14			10K SMD 0805 CHIP

R15			330E SMD 0805 CHIP

R16			1K  1% 0805 0.125W SMD

R17			620E SMD 0805 CHIP

R18*			10K SMD 0805 CHIP

R19			10K SMD 0805 CHIP

R20			100K 1% 0805 0.125W SMD

R21			2K4 SMD 0805 CHIP

R22			10K SMD 0805 CHIP

R23			2K7 1% 0805 0.125W SMD

R28			220E SMD 0805 CHIP

R29			20K SMD 0805 CHIP

R30*			1K8 0805 0.125W SMD

R31			220E SMD 0805 CHIP

R32			10E SMD 0805 CHIP

R33*			10K CHIP 1/8W  0805

NETWORK

RN1 TO RN9		10KE X4 x 1 COM 5 SMD

PRESETS

VR1			TRIMPOT 200E 10 TURN BOURNS

VR2			TRIMPOT 200E 10 TURN BOURNS

VR3			10K TRIMPOT HORZ

CAPACITORS

C1			10pF / 50V SMD CHIP 10% 0805

C2			10pF / 50V SMD CHIP 10% 0805

C3			10µF / 16V 2010MIL SMD ELE RAD

C4			100nF / 50V 10% 0805SM

C5			10µF / 16V 2010MIL SMD ELE RAD

C6			100nF / 50V 10% 0805SM

C7			100nF / 50V 10% 0805SM

C8			100nF / 50V 10% 0805SM

C9			220pF / 35V SMD CHIP 0805

C10			220pF / 35V SMD CHIP 0805

C12			18nF SMD 0805

C13			220pF / 35V SMD CHIP 0805

C14			220pF / 35V SMD CHIP 0805

PCB Components	5051A-MAIN-0306 IRD 551

=======================================================================

Ref Designator		Value

========================================================================

CAPACITORS

C16			18nF SMD 0805

C17			100nF / 50V 10% 0805SM

C18			10µF / 16V 2010MIL SMD ELE RAD

C19			33pF SMD CHIP 0805

C20			33pF SMD CHIP 0805

C21			100nF / 50V 10% 0805SM

C22			100nF / 50V 10% 0805SM

C27			3300µF / 35V 20% ELE RAD

C28			330µF / 25V ELE RAD

C29			10µF / 16V 2010MIL SMD ELE RAD

C30			100nF / 50V 10% 0805SM

C31			10µF / 16V 2010MIL SMD ELE RAD

C32			100nF / 50V 10% 0805SM

C33			10µF / 16V 2010MIL SMD ELE RAD

C34			100nF / 50V 10% 0805SM

C35			10µF / 16V 2010MIL SMD ELE RAD

C36			100nF / 50V 10% 0805SM

C37			1µF SMD CHIP 0805

C38			100nF / 50V 10% 0805SM

C39			100nF / 50V 10% 0805SM

C40			100nF / 50V 10% 0805SM

C41			100nF / 50V 10% 0805SM

C42			100nF / 50V 10% 0805SM

C43			ELE RAD 10µF, 16V

C44			100nf 50V CFI CHIP

CRYSTAL

X1			12MHz

X2			4.194304MHz

DIODE

D1 TO D4		1N4003 SMD MELF1

D5			RLS4148 T E11 SMD

ICs

U1			CS5460A 1PH POWER ENERGY 24 PIN TYPE SSOP

U2			89S52 UC 40 PIN 8 BIT FLASH 8KB  ( DIP40 )

U3			74HC4053 16 PIN 2 CHANNEL  ANALOG TYPE SO-16

U4			MNC 93C46N PI27 ATMEL MAKE ONLY

U5			LMV331 SMD 5 PIN LOW VOLTAGE  COMPARATOR SOT23-5

U7			TL431-CP ADJ PRECISION REG  SHUNT ( SOT23M1 )

U8			74HC14 SMD ( SO-14 ) HEX  INVERTING SCHMIT TRIGGER

U9			IC UA7805UC REG / LM7805CT +5V

U10			UA7905UC REG -5V FIX

U11, U12		OP07 8 PIN SMD OP AMP ( SO8 ) ULTRA LOW-OFFSET

CONNECTORS

J1			4 PIN 2.54mm MALE

J2			DUAL IN ROW UNLOC MALE STRAIGH  BURGE STRIP, 16 PIN

J3			3 PIN 2.54mm MALE

J4			4 PIN 2.54mm MALE

J5			DUAL IN ROW UNLOC MALE  STRAIGTH BURGE STRIP, 10 PIN

JP1, JP3, JP5	               3 PIN JUMPER  MALE BREG STRIP BLK

JP7			2 PIN 2.54mm JUMPER LINK SIL MJ-02

SWITCH

SW1			TACT SWITCH TINY

PCB Components	5051A-RS232-0706 IRD 570

=======================================================================

Ref Designator		Value

========================================================================

RESISTORS

R1			1K SMD CHIP 1% 0.125W  0805

R2			10K SMD CHIP 0805

R3			200E SMD CHIP 0805

R4			2.4K SMD CHIP 0805

R5			1K SMD CHIP 1% 0.125W  0805

R6			10K SMD CHIP 0805

R7			200E SMD CHIP 0805

R8			2.4K SMD CHIP 0805

R9			220E SMD CHIP 0805

CAPACITORS

C1			10µF / 16V 20%  ELE RAD

C2			10µF / 16V 20%  ELE RAD

C3			10µF / 16V 20%  ELE RAD

C4			10µF / 16V 20%  ELE RAD

C5			470µF / 35V 20% ELE RAD

C6			10µF / 16V 20%  ELE RAD

C7			100nF / 50V 10% 0805SM

C8			100nF / 50V 10% 0805SM

ICs

U1			TTL TO RS232 CONVERTOR 16 PIN

SMD SO-16

U2			OPTO ISOLATOR ISP817  4 PIN DIP

U3			OPTO ISOLATOR ISP817  4 PIN DIP

U4			UA7805UC REG/LM7805CT +5V

DIODES

D1 TO D4		IN4003 SMD MELF1

TRANSISTORS

Q1, Q2			MMBT3906 PNP SWITCHING SMD

SOT23

CONNECTORS

J1			3.96mm 3 PIN MALE ST MOLEX

J3			DUAL IN ROW UNLOC MALE

STRAIGHT BURGE STRIP, 10 PIN

J4			4 PIN ST MALE CONN JALEX

PCB Components	ZG3M2020MTSO

=======================================================================

Ref Designator		Value

========================================================================

CAPACITOR

C301, C302		CAP MP 1KV 47nF AXIAL

CORE

L301			FERRET  RING CORE 10 x 6 x 3

PCB Components	5051A-KBD-0306

=======================================================================

Ref Designator		Value

========================================================================

SWITCH

S1 TO S6		ITTD 6 KEY BOAR

DBLACK HEIGHT 6mm

CONNECTORS

J1			DUAL IN ROW UNLOC MALE

STRAIGHT BURGE STRIP, 16 PIN

PCB Components             LCD ASSY-16X2 HY-1602F-205 YELLOW

=======================================================================

Ref Designator		Value

========================================================================

CONNECTORS

J1			16 PIN ST BREG STRIP BLACK

MECHANICALS

=======================================================================

Ref Designator

========================================================================

20A MAINS CORD

FUSE 20A

FUSE HOLDER 20A

115V SOCKETS AMERICAN 16A  x 2

ON / OFF SWITCH 16A x 2

## GLOBAL SPECIALTIES

INSTRUMENTS

1486 highland ave,cheshire,CT. 06410

800-572-1028  www.globalspecialties.com

M-00374			       EDITION : 1  ( APRIL, 2007 )